# Error Factors for Enforcing a 160 µg/m3

# Total Carbon (EC+OC) Personal Exposure Limit

For purposes of enforcement, Total Carbon (TC) is defined as the sum of Elemental Carbon (EC) and Organic Carbon (OC).  Both EC and OC are measured by NIOSH Method 5040.

MSHA will issue a noncompliance citation when a valid personal TC exposure measurement demonstrates noncompliance with the 160TC µg/m3 PEL at a high level of confidence.  To this end, MSHA will continue adjusting each measurement according to an unexposed control filter (i.e., field blank) corresponding to each exposed filter.  As explained in more detail elsewhere, MSHA will further adjust the TC measurement, when necessary, to avoid interferences from extraneous sources of OC.  These adjustments are intended to minimize uncertainty due to potential systematic errors.

Neither of these adjustments, however, addresses uncertainty due to potential sampling and analytical errors (SAE).  Such errors, which reflect the imprecision inherent in any measurement process, cause individual measurements to deviate above or below the true value by a random but statistically quantifiable amount.  Measurement imprecision is quantified by the total coefficient of variation for SAE, or CVtotal.

As with all other single-sample M/NM noncompliance determinations, MSHA will address uncertainty due to SAE by allowing a margin of error before issuing a citation for exceeding the 160 µg/m3 TC limit.  This margin of error is designed to ensure that a citation will be issued only when a sample measurement demonstrates noncompliance with at least 95-percent confidence.  To achieve this 95-percent confidence level, the necessary margin of error is constructed by applying an error factor appropriate for the measurement being considered. The error factor is calculated as

EF=1+(1.645×CVtotal).

Each exposure measurement involves three components: the volume of air pumped through the filter that collects DPM, the deposit area of particles on the filter (cm2), and the laboratory analysis of EC and OC density within the deposit (µg/cm2).  CVtotal consists of three corresponding independent components — denoted CVP, CVD, and CVA — that respectively quantify the random variability associated with each of these factors.  To determine CVtotal for an individual EC or unadjusted TC measurement, the components of CVtotal are estimated separately and then combined according to a standard propagation of errors formula:

CVtotal=CVP2+CVD2+CVA2

Appendix 1 derives MSHA’s best estimates of the three CV components, based on the most recent laboratory data available.  Appendix 2, which has previously appeared on MSHA’s website in connection with earlier DPM error factors, explains details of the method by which updated estimates of CVA were calculated for EC and TC measurements.  Based on these estimates, CVtotal works out to be CVtotal[EC] = 5.7% for an EC measurement and CVtotal[TC] = 11.7% for an unadjusted TC measurement.

In accordance with Program Policy Letter No. P08-IV-1, some personal TC measurements will be ratio-adjusted in order to avoid interferences from extraneous sources of OC.  This will be done by multiplying the EC concentration measurement from the personal sample by the average (R) of TC-to-EC ratios taken from some number (N) of area samples.  For such a ratio-adjusted measurement,

CVtotal[TCadj]=CVtotal2[EC]+CVA2[R]

since (1) sampling and analytical measurement errors in the area samples are statistically independent of such errors in the personal sample and (2) airflow and deposit area errors cancel out when forming the ratio of TC to EC within each sample.  The last section of Appendix 1 derives a formula for computing CVA[R] for any given value of N.  Using these values, along with the value of CVtotal[EC] = 5.7% mentioned earlier, CVtotal[TCadj] ranges from 15.7% for N=1 to 7.4% for N=10.

In accordance with Program Policy Letter No. P08-IV-1, there are two situations in which MSHA will cite noncompliance:

- The 8-hour equivalent EC concentration measured by a valid personal sample exceeds 160 µg/m3 multiplied by the EF for EC measurements.  In this case,

EFEC = 1.095

Therefore, in this case, a citation will be issued when the personal 8-hour equivalent EC concentration measurement equals or exceeds 176 µg/m3.

- The 8-hour equivalent TC concentration (ECpersonal + OCpersonal) measured by a valid personal sample exceeds 160 µg/m3 multiplied by the EF for such measurements,

AND

the ratio-adjusted personal TC concentration exceeds 160 µg/m3 multiplied by the appropriate EF, which depends on N (i.e., the number of area samples used to form the adjustment).

In this case, EFTC = 1.192 and EFTC(adj) ranges from 1.259 for N=1 to 1.121 for N =10.  The general formula for EFTC(adj), for any given value of N, is:

EFTC(adj)=1+1.6450.057452+0.03419222+0.14562N

Therefore, in this case, a citation will be issued when:

  - the personal 8-hour equivalent TC concentration measurement equals or exceeds 191 µg/m3, and

  - the ratio-adjusted 8-hour equivalent TC concentration equals or exceeds the value tabulated below.

`

| Number of Area Samples Included in Average Ratio† | Citation Threshold for TCadj |
| --- | --- |
| 1 | 202 |
| 2 | 192 |
| 3 | 187 |
| 4 | 185 |
| 5 | 183 |
| 6 | 182 |
| 7 | 182 |
| 8 | 181 |
| 9 | 180 |
| 10 | 180 |

†To be included in the average TC-to-EC ratio, the EC content of an area sample (i.e., the denominator of the ratio) must be at least 5 µg/cm2.

We believe that our estimates of CVtotal, and the error factors they produce, are based on the best scientific data currently available and adequately reflect current sampling and analytical errors.  However, we recognize that future improvements in sampling and/or analytical technology may reduce the random variability associated with measuring carbon concentrations.  Therefore, MSHA may update these error factors based on future experimental data.

# Appendix 1.

# Components of CVtotal

CVP: Variability in volume of air pumped through the filter

Variability in air volume depends on three factors: (1) variability in the initial setting of the pump rotameter to a calibration mark when sampling begins, (2) pump calibration errors, and (3) variability in air flow during the sampling period.  Based on Bowman et al. (1984), MSHA estimates that uncertainty due to the combined effects of calibration errors and flow rate variability is represented by a coefficient of variation (CV) no greater than 3%.  Based on the experimental results described by Tomb (1994), MSHA estimates that the CV component associated with variability in setting the rotameter ball is approximately 3%.  Since variability in the initial flow rate is independent of calibration of the pump rotameter and variability in flow rate during sampling, these two uncertainty components can be combined as follows to yield the CV representing uncertainty in total volume of air pumped:

CVD = (0.03)2+(0.03)2=0.042

or 4.2 percent.

CVD: Variability in area of dust deposited on filter

Variability in SKC sampler performance is manifested as variability in the area and uniformity, or density, of the particulate matter deposited on the filter.  Variability in the density of the deposit is included in the estimated value of CVA and is discussed below.  Variability in the deposit area is addressed by CVD.

Since 2001, the manufacturer of the sampling device (SKC, Inc. Eighty Four, PA.) has made a number of improvements designed to reduce variability in the deposit area.  MSHA’s current estimate of CVD is based on a 2005 study assessing the performance of the improved device.  Using data collected on 94 improved SKC samplers, that study found the coefficient of variation in deposit area to be 0.005 (i.e., 0.5 percent).  Therefore,

CVD = 0.005

or 0.5 percent.

CVA: Analytical measurement imprecision

Analytical measurement imprecision refers to the random variability of repeated measurements of an analyte, performed on different punches taken from the same filter, within the same or different laboratories.  In addition to imprecision in the instrumentation, this encompasses random variability in the punch area and in the density of the deposit, but not in the deposit area.  Variability in the deposit area (a form of sampling variability) is quantified separately as CVD above.

To estimate CVA, MSHA combined data obtained from two sources.  Inter-laboratory variability was estimated, as in previous versions of the DPM error factor, based on the original “paired punch comparison” carried out as part of the “31-Mine Study.” The estimate of intra-laboratory variability, however, was updated to reflect current performance within the MSHA laboratory that analyzes MSHA’s DPM samples.  This updated estimate of intra-laboratory variability was based on 364 paired punches (i.e., two punches from the same sample filter) analyzed from January 3, 2006 through April 3, 2008.  The independent estimates of inter- and intra-laboratory measurement variability were then combined to form an estimate of overall CVA.  For both component estimates, a square-root transformation was applied to each TC(OC+EC) and EC measurement to stabilize the variance in the statistical analysis.  Appendix 2 contains a justification for using this transformation and explains how it can be used to estimate intra-laboratory CVA as a function of the filter loading.

Inter-laboratory Variability

In the original paired punch comparison, 621 filters were analyzed using two standard punches taken from each filter.  One punch (labeled “A”) was always analyzed in MSHA’s laboratory.  The second punch from the same filter (labeled “B”) was either analyzed in MSHA’s laboratory or in one of three other laboratories.

A repeated measures, random effects Analysis of Variance (ANOVA) was performed to derive an estimates of the inter-laboratory components of analytic measurement imprecision, based on the available data from all four laboratories.

For reasons explained in Appendix 2, the model used in the ANOVA was:

XjA−XijB=Δi+λij(i≠MSHA)+εij

where

X is the analytic result from punch A or B reported in g/cm2;

i indexes the laboratory analyzing Punch B;

j indexes a specific filter;

i is a fixed effect, representing the systematic difference between MSHA’s punch A results and the punch B results at laboratory i;

ij is a random, Normally distributed, inter-laboratory effect with mean = 0 and variance = σλ2;

ij is a random, Normally distributed, intra-laboratory error with mean = 0 and variance = σε2.

Inter-laboratory imprecision in this model was quantified by σλ2 and, based on the ANOVA, estimated to be σλ2=0.01642 for EC and σλ2=0.07713 for TC.

Intra-laboratory Variability

To estimate intra-laboratory imprecision, represented by σε2, the method of Appendix 2 was applied to MSHA’s new body of paired punch data, compiled from January 3, 2006 through April 3, 2008.  Accordingly, σε2 was estimated by the squared sample standard deviation of X1−X2, using the 364 available paired punches from that period.  The resulting estimates areσε2= 0.00768 for EC and σε2= 0.01412 for TC.

Combined Analytical Uncertainty

The composite estimate of analytical measurement uncertainty, including both intra- and inter-laboratory imprecision, is represented by σT2=σλ2+σε2.  Therefore,

σT=σλ2+σε2=0.15524 for EC

Appendix 2 shows that for a carbon measurement (X) based on a single punch, the coefficient of variation in analytical error is

CVμ[X]=σT2μ

where  is the true carbon loading (g/cm2) on the filter.  However, to reduce analytical measurement uncertainty, MSHA routinely averages the results (X1 and X2) from two punches taken from an exposed filter and then subtracts the corresponding result (B) from an unexposed control filter.  The adjusted measurement based on averaging X1 and X2 can be expressed as

Y=X1+X22−B.

To simplify the notation in what follows,   will be used to represent  .  As shown in Appendix 2, Var[Xi] (i.e., the variance of Xi) is 22.  Similarly, Var[B] = 22E[B], where E[B] is the “expected” or mean density of carbon measured on a control filter.  Therefore, assuming independent analytical measurement errors for X1, X2, and B,

Var[Y]=1222 Var[X] + Var[B]σ2μ + 2σ2E[B]σ2μ+2E[B]σ2E[Y]+E[B]+2E[B] since E[Y]  is  μ−E[B]σ2E[Y]+3E[B]

It follows that:

CVA=CV[Y]Var[Y]E[Y]σE[Y]+3E[B]E[Y]

Based on data compiled from MSHA’s analysis of carbon measurements on available control filters between December 4, 2002 and April 3,2008,

EB=0.1  μg/m2 for EC .

Furthermore, using an airflow rate of 1.7 L/min, a deposit area of 8.0425 cm2, and a nominal sampling duration of 480 min:

E[Y] ≥ 16.23 g/cm2

for EC or TC 8-hour equivalent concentration levels at or above 160 µg/m3.  Therefore, substituting the respective EC and TC values of σT for   in the formula for CVA and noting that CVA decreases as E[Y] increases, it is evident that:

CVA≤0.0388841  or 3.89 percent for EC

CVA[R]: Analytical measurement imprecision for an average ratio adjustment factor

Suppose N samples are used to determine the average ratio of TC to EC in a mine.  The TC-to-EC ratio (R) exhibited by a single given sample is calculated as follows:

R=TC1+TC22−BTC EC1+EC22−BEC

where the two analytical measurements shown for both EC and TC are all made on the same exposed filter.   BEC and BTC are measurements made on the same unexposed control filter, applied as corrections for the EC and TC measurements respectively.  R, the ratio of TC to EC used to adjust a personal TC exposure measurement, is defined as the sample average of N such ratios.

Within each sample used to estimate CVA[R], let R1=TC1EC1 and R2=TC2EC2.  Let W=LogeR1R2.  Then, assuming measurement errors in R1 and R2 are identically distributed and statistically independent, Var[W] ≈ 2×CV2[Ri] for i = 1 or 2.  Utilizing this relationship, MSHA’s most recent body of paired punch data (January 3, 2006 through April 3, 2008) was used to estimate CV[Ri].

Figure 1 plots W against the average EC measurement from a single filter and shows that variability in W increases markedly for EC loadings less than 5 µg/cm2.  Therefore, the estimate of Var[W] presented here is restricted to samples for which the average EC ≥ 5 µg/cm2.  Based on the sample standard deviation calculated from the 326 (out of 364) measurement pairs meeting this criterion, the estimated value of Var[W] is (0.03419)2.  Therefore, the estimated value of CV2[Ri] is (0.03419)2/2.

Figure 1.  Natural logarithm of R1/R2, plotted against the average EC measurement from two punches of the same filter, based on 364 paired punches analyzed from January 3, 2006 through April 3, 2008.

As indicated above, E[BEC] = 0.1 µg/cm2 whereas E[BTC] = 6.0 µg/cm2.  Therefore, for filters in which EC ≥ 5 µg/cm2, BEC has a negligible effect on the calculation of R.  Consequently, the contribution toward CVA[R]of random error in the control filter adjustment can be considered to be based entirely on BTC.

To obtain a conservative estimate of this contribution, the sample standard deviation of  (R1 − R2) was calculated for just those samples in which TC ≤ 10 µg/cm2 and then divided by 2.  Based on the 15 samples (out of 364) for which TC ≤ 10 µg/cm2, the contribution of error in a single control filter adjustment toward overall analytical variability inR is conservatively estimated to be 0.205942, or 14.6 percent.

To construct an estimate of CVA[R], the estimated value of CV2[Ri] is divided by two, to reflect the fact that each value of TC and EC used to form a ratio is based on the average of two independent punch measurements.  The result is then combined, by means of a standard propagation of error formula, with the squared estimated contribution of the control filter correction for TC.  SinceRrepresents the average of N independently determined ratios, CVA2[R] is further reduced by a factor of N.  This yields the following estimate, expressed as a decimal fraction:

CVAR = 0.03419222+0.14562N.

If the average TC from two unexposed control filter punches, instead of from one, were used to correct each TC measurement, then 0.1456 would be replaced by 0.1030 in this formula.

For reasons explained earlier, the estimated value of CVA[R]is valid only when the samples used to calculateRall exhibit average EC loadings ≥ 5 µg/cm2.

# Appendix 2.

# Use of Variance-Stabilizing Transformation for

# Analysis of TC(OC+EC) and EC Measurement Variability

Let i index a specific filter, and let Xi1 and Xi2 denote two carbon measurements (g/cm2) made using two punches from that filter.  As noted in the documentation for NIOSH Method 5040, the variance of a carbon measurement made using this method (Var[Xi]) is roughly proportional to the carbon loading (µg/cm2) on a filter.  This relationship can be expressed as

Var[Xi]=λ2μi						(Eq. 1)

where λ2 is a constant and µi is the true loading on the ith filter.  Since µ varies but λ is constant, it follows that the coefficient of variation (CVμ[X]), which quantifies measurement variability relative to any given loading, decreases as µ increases:

CVμ[X]=Var[X]μ=λμμ=λμ				(Eq. 2)

To estimate λ, and thereby to calculate CVµ[X] as a function of filter loading, a variance-stabilizing square-root transformation was applied to each measurement.  Using the standard propagation of error formula applicable to Eq. 1,

VarXi≈λ24

for a carbon measurement at any filter loading.  Based on this approximation, and assuming independent measurement errors in Xi1 and Xi2,

VarXi1−Xi2=2×VarXi=λ22			(Eq. 3)

Consequently,

λ=2×Var Xi1−Xi2 12 =σ2				(Eq. 4)

where σ denotes the standard deviation of the differencesXi1−Xi2.

From Equations 1 and 4 it follows that Var[Xi]=2σ2μi, and combining Equations 2 and 4 yields the formula used to quantify EC or TC measurement variability at a given filter loading:

NIOSH Manual of Analytical Methods, Fourth Edition, at www.cdc.gov/Niosh/nmam/pdfs/5040.pdf

CVtotal represents the magnitude of probable sampling and analytical errors.  Sometimes called the relative standard deviation (RSD), CVtotal is defined as the ratio of the standard deviation of sampling and analytical measurement errors to the true value of whatever quantity is being measured.  It can be  expressed either as a fraction (e.g., 0.1) or as a percent (e.g., 10 percent) of the true value.

This measurement is based on the average EC from two punches of an exposed filter, reduced by the EC from a single punch of an unexposed control filter.

This refers to a TC measurement that has not been adjusted using the ratio of TC to EC in any area samples.  However, like the EC measurement, it is based on the average TC from two punches of an exposed filter, reduced by the TC from a single punch of an unexposed control filter.

To be included in the average ratio, the measured EC content of an area sample must be at least 5 µg/cm2.

Each area sample ratio is based on averaging TC and EC from two punches of an exposed filter, reduced by the corresponding results from a single punch of an unexposed control filter.  As discussed in Appendix 2, CVtotal for the ratio-adjusted TC measurements (and consequently the error factors derived from them) could be substantially reduced if MSHA based its blank filter correction for area TC measurements on the average TC density (µg/cm2) detected on two punches from the associated unexposed control filter.

Bowman et al. (1984), Precision of Coal Mine Dust Sampling, CDC (NIOSH); NTIS No. PB-85-220-721

Tomb (1994) Memorandum dated Sept. 1 to Chief, Division of Health, CMS&H, MSHA, Subject: Determination of the Precision of Setting the Rotameter Ball to a Calibration Mark on Personal Respirable Dust Sampling Pumps. (available from MSHA’s CMS&H single-sample rulemaking record)

Noll, J.D., R.T. Timko, L. McWilliams, P. Hall, and R. Haney. “Sampling Results of the Improved SKC Diesel Particulate Matter Cassette,” Journal of Occupational and Environmental Hygiene, 2:29-37, January 2005.

MSHA’s Report on Data Collected During a Joint MSHA/Industry Study of DPM Levels  in Underground Metal and Nonmetal Mines (January 6, 2003), available in the M/NM DPM Rulemaking Record (RIN 1219-AB29).   The report contains a full description of the paired punch comparison, along with all of the data collected in connection with the study.  Although the study mainly addressed variability in TC measurements, EC was separately measured in the course of the laboratory analysis of each punch.  Consequently, both the TC and the EC analytic results reported in connection with this study are in the DPM rulemaking record.

Because of the particular experimental design employed, the results combine purely analytical imprecision with variability in the density of the particulate deposited on the filter and with variability in the way the two punches were handled prior to analysis.  Therefore, the estimate of CVA presented here covers all three of these uncertainty components.

Loge denotes the natural logarithm (i.e., the logarithm to the base e).

This is because Var[Loge(Ri)] ≈ Var[Ri]  E2[Ri].  See Ku, H.H. “Notes on the Use of Propagation of Error Formulas”, Precision Measurement and Calibration, NBS Special Publication 300, Vol. 1, 1969.  pp. 331-341. Note also that Log(R1/R2) = Log(R1) – Log(R2).  Therefore, under the stated assumptions, E[R1] = E[R2] and Var[Log(R1/R2] = 2×Var[Log(Ri)].

Since EC appears in the denominator of each ratio, even relatively small errors in the EC measurement can have a large impact on the ratio when EC < 5 µg/cm2.  The difference in Var[W] for EC above and below this threshold is statistically significant.

The factor of 2appears because Var[R1 − R2] = 2σR2, where σR is the standard deviation of Ri.  To calculate the CV component representing the actual contribution, σR must be divided by the ratio of TC to EC, which is unknown and anticipated to vary from mine to mine.  However, since this ratio is always ≥ 1, the CV component is necessarily ≤ σR.  In this sense, σR  provides a conservative estimate.

If the average from two unexposed control filter punches were used to adjust each TC measurement, instead of from one punch, then  σR would be reduced by an additional factor of 2.  In this case, the conservatively estimated contribution to CVA[R]would be reduced from no more than 14.6% down to no more than 10.3%.

NIOSH. Manual of Analytical Methods (NMAM). Chapter Q,  “Monitoring of Diesel Particulate Exhaust in the Workplace,” Third Supplement to NMAM, 4th Edition (2003), p. 234.

Ku, H.H. “Notes on the Use of Propagation of Error Formulas”, Precision Measurement and Calibration, NBS Special Publication 300, Vol. 1, 1969.  pp. 331-341.